| Equipment Name: | Cooling Pad |
| --- | --- |
| Required for: | Placing slides for testing |
| Equipment ID: | CP01 |
| Document No: | CPDQ01 |

Table of Content

| Sr. No | Description | Page No. |
| --- | --- | --- |
| 1 | Document Approval | 2 |
| 2 | Design Qualification Summary | 2 |
| 3 | Basis of Design | 3 |
| 4 | System Description | 3 |
| 5 | Technical Specification of Component | 4 |
| 6 | Accompanied Documentation | 5 |
| 7 | Conclusion | 5 |

# DOCUMENT APPROVAL:

| Details | Designation | Signature | Date |
| --- | --- | --- | --- |
| Prepared By |  |  |  |
| Checked By |  |  |  |
| Approved By |  |  |  |

DESIGN AND QUALIFICATION SUMMERY:

  - Material Safety: Use of bio-compatible anodized aluminium and chemical-resistant PLA housings.

  - Control Accuracy: Closed-loop PID control capable of stabilizing temperature bands to ±0.25°C.

  - Electrical Safety: Noise isolation, 100-ohm current-limiting loops, and fail-safe overcurrent margins. OBJECTIVE AND SCOPE

## OBJECTIVE & SCOPE

Objective: To provide documented verification that the engineering design of the Peltier Cooling Pad is compliant with user parameters, bioscience laboratory specifications, and global safety guidelines. This ensures precise, software-controlled, solid-state heating and cooling of laboratory substrates (0.0°C to 50.0°C).

Thermal Loop: 4x TEC1-12705 Peltier grid array clamped under liquid-jacket heatsinks.

Control Subsystem: ESP32 SoC microcontroller executing real-time Proportional-Integral-Derivative (PID) software, connected to a BTS7960 high-current driver and 100k NTC temperature feedback.

Liquid Cooling: 12V DC water pump and water blocks for hot-side heat dissipation.

User Interface: Local 0.91" I2C OLED display and web server hosted through the controller itself.

# BASIS OF DESIGN:

Peltier Grid Sizing: A grid of 4x TEC1-12705 tiles in series-parallel configuration. At 12VDC, each tile draws ~3.5A, culminating in a 14A peak. The BTS7960 driver (rated up to 43A) comfortably manages this load.

Power Supply Over-Rating Margin: A Mean Well 12V 30A (360W) SMPS is

integrated. The peak draw of the peltiers (168W) + water pump (4.8W) + exhaust fan (2.1W) + logic (2.5W) totals 177.4W. This provides a safety margin of over 100%, preventing supply-line brownouts.

Thermal Feedback Loop: Standard 100k NTC temperature sensors embedded in the anodized black plate read raw analog voltages, which the ESP32 smooths via a 10-point rolling filter.

Active Control Rates: Custom PID parameters interpolated dynamically via a calibration map ensure rapid slew rates:

Cooling Rate: ≥1.5°C/min (from room temperature to 5.0°C)

Heating Rate: ≥2.5°C/min (from room temperature to 37.0°C)

- SYSTEM DESCRIPTION:

Physical Layout: Features a compact 140x160x125mm housing printed in high-grade PLA. The upper surface comprises an anodized black aluminium active cooling pad. Beneath this plate, dual liquid blocks circulate water via brushless 12V pump to carry hot-side heat away.

Control Electronics: Guided by the ESP32 An I2C 0.91-inch OLED screen provides instant local readouts of Setpoint, Live Temp, SSID, and IP address.

Embedded Firmware: Hosts an interactive webserver portal. Users can configure multi-step cycling profiles, adjust PID tuning parameters, download raw log files, and adjust calibration tables via any standard laboratory web browser.

Electrical Safety: Integrates high-conductive 14 AWG silicone power cables, 100-ohm current limiters, and dedicated software watchdogs to kill power immediately if thermistor signals are lost.

# TECHNICAL SPECIFICATIONS OF COMPONENTS:

| Component Item | Technical Details | Manufacturer |
| --- | --- | --- |
| ESP32 | Main controller | Espressif Systems |
| BTS7960 | 52A High-capacity motor driver | Infineon Technologies |
| TEC1-12705 | 40x40mm, 12V 5A peltiers | HB Electronic |
| 100k NTC Thermistor | range -40°C to +150°C | Murata |
| 0.91-inch OLED | SSD1306 driver, 128x32 pixels, blue color | Solomon Systech |
| Water Pump | 12V submersible DC water pump | Local manufacturer |
| DC Fan | 40X40mm DC fan for cooling | Local manufacturer |
| Power supply | 12V 30A, SMPS for AC-DC | Maxicom |

# 6. ACCOMPANIED DOCUMENTATION:

General Assembly: Outlines PLA tolerances, clamping forces, and fluid-line routing.

Electrical Schematic: Details pin-to-pin mapping of the ESP32, BTS7960, OLED, and NTC sensor.

Embedded Firmware Source Code: Version-controlled file including library references.

Operations & Maintenance Manual: Procedures for WiFi setup, profile execution, and liquid cooling reservoir flushes.

Component Datasheets: Official datasheets for the SMPS, TEC1 Peltier modules, and PLA filaments.

NTC Calibration Log Sheet: Calibration curve parameters ($B=3950$) signed by the lead test engineer.

7. CONCLUSION:

The Design Qualification study concludes that the mechanical, electrical, thermal, and software layout of the Peltier Cooling Pad fully satisfies the parameters defined in the parent User Requirement Specification (URS).

The physical component selections provide ample safety margins (over 100% on power feed), reliable closed-loop PID control accuracy (±0.25°C), and rigorous safety fail-safes (automatic thermal shutdown on sensor or water-loop failures).

Therefore, the Design Qualification (DQ) phase of this system is officially APPROVED. The equipment is cleared to proceed to fabrication, wiring assembly, and subsequent site installation and operational qualification phases.