| Equipment Name: | Cooling Pad |
| --- | --- |
| Required for: | Placing slides for testing |
| Equipment ID: | CP01 |
| Document No: | CPIQ01 |

Table of Content

| Sr. No | Description | Page No. |
| --- | --- | --- |
| 1 | Document Approval | 2 |
| 2 | Objective | 2 |
| 3 | Scope | 3 |
| 4 | System Identification and General Description | 3 |
| 5 | Acceptance Criteria | 4 |
| 6 | Installation Checklist | 4 |
| 7 | Major Component Validation | 5 |
| 8 | Utilities and Services Required | 5 |
| 9 | Recommended Parts and Lubricants | 6 |
| 11 | Non-Conformance and Corrective Action Log | 6 |
| 12 | Certification Report | 6 |
| 13 | Report Approval | 7 |
| 14 | Revision History | 7 |

| Details | Designation | Signature | Date |
| --- | --- | --- | --- |
| Prepared By |  |  |  |
| Checked By |  |  |  |
| Approved By |  |  |  |

- DOCUMENT APPROVAL:

- OBJECTIVE:

  - The objective of this Installation Qualification (IQ) is to provide formal, documented verification that the Peltier Cooling Pad has been physically delivered, unboxed, and installed in the target laboratory site according to manufacturer specifications and relevant safety and quality standards.

  - This qualification verifies that:

- The physical hardware has been received intact without shipping damage.

- The mechanical, fluidic, and electrical connections are safely established.

- The installation site meets all environmental, ventilation, and power utility guidelines.

- SCOPE:

  - This Installation Qualification protocol covers the physical boundaries and site utilities associated with the installation of the Peltier Cooling Pad.

  - Physical Equipment: Dual-compartment PLA enclosure housing, 4-Peltier solid-state grid, anodized active cooling plate, brushless DC water pump, and 12V SMPS.

  - Site Interfaces: Mains AC socket (110/220V), bench spacing (140x160mm minimum footprint), and demineralized cooling water reservoir supply.

  - Logic Connections: Local 0.91" I2C OLED display lines and ESP32 wireless telemetry network.

- SYSTEM IDENTIFICATION & DESCRIPTION:

  - Equipment Name: Peltier Cooling Pad

  - Equipment Model/Part Number: CP01

  - Serial Number:  01

  - Manufacturer: Candore Bioscience

  - Date of Manufacture: 10/06/2026

  - Installation Date: 10/06/2026

  - Installation Location: Santej, Ahmedabad

| System | Components |
| --- | --- |
| Control System | ESP32 S3 microcontroller |
| Display Unit | 0.91-inch OLED |
| Motor Driver | BTS7960, 43A motor driver |
| Temperature Sensing | 10K NTC Thermistor |
| Cooling System | 40x40mm TEC1-12705, Aluminium water block, 12V DC water pump |
| Power Supply | 12V, 30A SMPS |
| User Interface | Web based UI using local host wifi |

- ACCEPTANCE CRITERIA:

  - To achieve passing status and qualify the physical installation, the following absolute criteria must be met without deviation

  - Equipment Integrity: No visible physical damage, chips, or fractures on the 3D-printed PLA casing, active pad, or fluid lines.

  - Component Completeness: All bill-of-materials components are physically present.

  - Fluid Safety: Closed-loop water lines must show zero leakage or condensation seepage when pressurized by the 12V DC water pump.

  - Electrical Insulation: High-power 12V DC lines must be fully isolated from low-voltage 3.3V/5V logic wires.

- INSTALLATION CHECK LIST:

  - Chassis Assembly: Visual inspection of PLA casing; verify 0.2mm tolerance joints.

  - Active Pad Mounting: Verify the white surface of the peltier is flat and clean.

  - Tubing Connections: Check whether the silicone pipes are properly connected to each end of the water block and pump.

  - High Current Wiring: Verify the 12V power cables are clamped properly inside SMPS.

  - Sensor Placement: The NTC thermistor should be properly mounted on the surface of peltier.

  - OLED Display: The 0.91-inch OLED display should be properly fixed in the front window.

  - Placement: The device should be placed on a flat surface that does not vibrate.

- MAJOR COMPONENT VALIDATION:

  - ESP32 Microcontroller: Checked for secure header pins mounting and successful boot sequences.

  - BTS7960 Driver Module: Checked for 12V wires properly connected in the terminals.

  - TEC1-12705: Properly placed on aluminium water blocks.

  - 12V SMPS: Checked for proper wiring and ventilation.

  - Brushless Water Pump: Confirmed secure pipe connection and proper mounting beside the main body of the device.

  - Exhaust Fan: Verified the airflow and direction of the 40x40mm DC fan.

- UTILITIES & SERVICES REQUIRED:

  - AC Electrical power, Single-phase AC 220V ±10%.

  - Surge-protected wall socket, minimum 5A line rating.

  - Water should be filtered, don’t use tap water directly.

  - Requirement: Safe level surface, rear space >= 10 cm for fan exhaust.

  - Standard 2.4GHz Wi-Fi for remote webserver control and monitoring.

- RECOMMENDED SPARE PARTS:

  - TEC1-12706 Thermoelectric Peltier.

  - 10K NTC Thermistor.

  - 12V DC water pump.

  - HM-501 Thermal grease.

- NON-CONFORMANCE & CORRECTIVE ACTION:

| Date, Sign | Remark |
| --- | --- |

- CERTIFICATION REPORT:

  - All physical parts and major electronic sub-components are present, undamaged, and correctly assembled.

  - Site utilities (AC electrical mains, spatial clearances, and fluid cooling reservoir parameters) meet all specification thresholds.

  - Active safety checks, ground lines wiring, and non-conformance log deviations are verified closed.

- APPROVAL REPORT:

| Details | Designation | Signature | Date |
| --- | --- | --- | --- |
| Prepared By |  |  |  |
| Reviewed By |  |  |  |
| Approved By |  |  |  |

- REVISION HISTORY:

| Revision No. | Change Control No. | Reason for revision | Changes Incorporated |
| --- | --- | --- | --- |